## Slide 1: Photon Detector for Compton Polarimeter at SuperKEKB
D. Charlet et al., 2023 JINST 18 P10014

## Slide 2: Introduction
  Photon Detector
  Requirements :
    Rate : 250 MHz
    Independent Bunch-by-Bunch Measurement
    Dynamic Range : from 100 MeV to 10 GeV
    Radiation Hard : 0,6 MGy for 6 months of operation
    Precise : 1 % of statistical uncertainty every 5 minutes
Detector Location

## Slide 3: Acquisition Chain
Scintillator
PMT
FPGA
IDROGEN board
White Rabbit node
Analog Front-end
+ ADC (>=10bits)
Every 4ns
2500 histograms
Fast treatment: fit
Accelerator RF
1rst Bunch identifier
To remote storage (10MB/min)
Reduced data
To online monitoring
clock
Update program, constants
UV filter
Clock to Laser Synchronization

## Slide 4: Scintillator Choice : Requirements
  Requirements :
    Radiation Hard : 0,6 MGy for 6 months of operation
    Very fast Scintillation : 4ns between bunches
  BaF2:
    Sub-ns scintillation
    Density higher than organic scintillators
    Not hygroscopic
    Fast / Slow components
    Filter for the slow component

## Slide 5: Scintillator Choice : Undoped BaF2
  BaF2 scintillator :
    = 1800 ɣ/MeV (fast)
    = 10 000 ɣ/MeV (slow)
  Datasheet Saint-Gobain BaF2

## Slide 6: Scintillator Choice : BaF2:Y
  EPIC Crystal Scintillator (BaF2:Y)
    = 3 040 ɣ/MeV (fast)
    = 12 160 ɣ/MeV (slow)
  Likely a typo in the datasheet on the Light Yield
  Ordered
  Datasheet Epic Crystal BaF2:Y

## Slide 7: UV Filter Transmission
Datasheet Knight Optical Filter
BaF2:Y
BaF2

## Slide 8: Slow Component Filtering
  Fraction of the slow component after filtering :
    For BaF2 Scintillator : 0.16 %
    For BaF2:Y Scintillator : 0.08%
  Light output :
    For BaF2 Scintillator :
    For BaF2:Y Scintillator :
  Measurements of the Filter Transmission and Light Yields to be made

## Slide 9: Photomultiplier Tube
Requirements : Fast Response (Low Rise Time and Low Transit time Spread) and not blind to UV.
Hamamatsu Photomultiplier R3377 ( Datasheet Hamamatsu R3377 )
  Nominal Gain :
  Silica Glass Window : Respond to UV.
  Linear Focusing Dynodes : fast response
Typical Spectral Response
 | Nominal Gain | G = 104
Rise Time (ns) | 0.80 | 1.39
Transit Time Spread (ns) | 0.37 | 0.64

## Slide 10: Photomultiplier Tube
Requirements : Fast Response (Low Rise Time and Low Transit time Spread) and not blind to UV.
Hamamatsu Photomultiplier R3377 ( Datasheet Hamamatsu R3377 )
  Nominal Gain :
  Silica Glass Window : Respond to UV.
  Linear Focusing Dynodes : fast response
Time Characteristics vs. Supply Voltage
  Hamamatsu Handbook
 | Nominal Gain | G = 104
Rise Time (ns) | 0.80 | 1.39
Transit Time Spread (ns) | 0.37 | 0.64

## Slide 11: Expected Sensitivity
Hypothesis : Light Collection = 10 %, no light attenuation yet, Gpmt = 104
We need to lower the current at the anode maybe by lowering the light flux
 | BaF2 | BaF2:Y | Max. Ratings
 |  |  | 96
() |  |  | 
Ipeak,anode (mA) |  |  | 
 | BaF2 | BaF2:Y | Max. Ratings
 |  |  | 96
Ipeak,anode (mA) |  |  | 

## Slide 12: Solutions
Concessions to be made on time and/or energy resolution
  Reduce the sampling rate
  Lower the light intensity at the photocathode with a diaphragm or another filter

## Slide 13: Limitations
Linearity
  Constraint on the mean and peak anode current ( and
Time Response
  Transit Time Spread and most characteristics times scale as
Energy Resolution
  Electronic Noise scales as and Photocathode Energy Resolution as
We need a trade off between pile up (time response), energy resolution, linearity and PMT ageing

## Slide 14: Simulations
Goal : Find the maximum energy resolution that is acceptable
Monte Carlo Simulation
  Energy spectrum of the scattered photons
  Fitting of energy spectrum
Acquisition Chain Simulation
  Scintillator and PMT response
  Electronic noise and Time response
  Filter
  ADC

## Slide 15: Pulses Simulation
D. Charlet et al., 2023 JINST 18 P10014
First Approach :
Generate 106 pulses ( 5 minutes ) with the Compton energy distribution.
Simulate the acquisition chain.
Export the reconstructed energy to fit the data.
But the acquisition chain cannot simulate 106 pulses (Memory Limits)

## Slide 16: Acquisition Chain - Energy Resolution
Other Approach
  Energy sampling on the range 0.9 to 10 GeV
  Simulation of 200 pulses with a fixed energy
  Determination of the energy resolution
  Monte Carlo simulation with this energy resolution
  Fit of the simulated data
What is in the Simulation
  Poisson Distribution
  Excess Noise factor
  ADC

## Slide 17: Simulation Parameters
Rate = 15 MHz
Sampling Rate = 3 GHz
Nbits = 10 bits
ADC Range = 1,5 V
𝜏fast = 0,8 ns
Excess Noise Factor : 1.2

## Slide 18: Results on Energy Resolution
Notes: A = 10.815 ± 0.137 %·\sqrt{GeV} | B = 5.000e-19 ± 0.000e+00 %

## Slide 19: What’s Next ?
Measurements on the Scintillators and Filter → Conclude on the expected Light Yield
Light Collection Measurement
Choose a PMT Gain and how much we want to reduce the light flux

## Slide 20: Back up – Gain vs. Supply Voltage

## Slide 21: Back up – Time Response vs. Vsupply (R3377)

## Slide 22: Back up – Voltage Divider Constraint

## Slide 23: Back up – Voltage Divider Constraint
and goes from 1 to 3 kV
at G = 104
Voltage divider given with the PMT cannot be changed